## Slide 1
BacT/ALERT® 3D Field Service Engineer Training
Lesson 4 System Installation
        4.1 BacT/LINK Overview
        4.2 LIS & Simulator Overview
        4.3 Installation Procedure
        4.4 Preventative Maintenance

## Slide 2
BacT/ALERT® 3D Field Service Engineer Training
Lesson 4 System Installation
Chapter 4.1 BacT/LINK Overview

## Slide 3: Objectives
Discover why and how the BacT/LINK is utilized by our customers
Understand the basic principles of BacT/LINK
Know how to setup & configure BacT/LINK
Know how to test BacT/LINK connections
Know how to use the LIS Simulator Software

## Slide 4
What is BacT/LINK
What does BacT/LINK provide the Customer?
  Empowers BacT/ALERT system to receive data from and to transmit to the Host (LIS) System
  Complies with standards as defined by the Clinical Laboratory Standard Institute, or CLIS, (originally written by the American Society for Testing and Materials or ASTM)
  It is able to perform both manual and automatic communication sessions between the Host and the Instrument or DBMS

## Slide 5
Configuration
How Does BacT/LINK appear in the Clinical situation?
SelectLinkTM
SignatureTM
L I S
L I S

## Slide 6
LIS Specifications
What are the RESOURCE DOCUMENTS available to us
  LIS1-A : Specification for the Transfer of Messages Between Clinical Laboratory Instruments and Computer Systems – Low Level Protocol (originally referred to as ASTM E 1381-91)
    How we physically effect the transfer
  LIS2-A : Standard Specifications for Transferring Information Between Clinical Instruments and Computer Systems – High Level Protocol (originally referred to as ASTM E 1394-91)
    How we order our codes and what codes do we use
  BacT/LINK External Specifications ( April 2003)
    This is our instruction manual from bioMérieux

## Slide 7
Low Level Protocol
PHYSICAL LAYER (section 5.x) – sets standards for electrical and mechanical connection between Instruments and LIS
  Defines data bits – 7 or 8
  Defines Parity – Even, Odd, None
  Defines stop bits – 1 or 2
  Sets baud rate (speed) – 1200 2400 4800 9600 (preferred)
  Sets specifications for cables and connectors –
  25 pin connector – pins 1 2 3 & 7 used
  Sets specifications for cables and connectors –
  9 pin connector – pins 1 2 3 & 5 used

## Slide 8
Low Level Protocol
Data Link Layer (section 6.x) – set ground rules for communications
  Information flows in one direction at a time
  Replies are required after information is sent
  Two systems actively operate to transfer information
  Neutral state – neither system sending or receiving information
  Three distinct phases for transferring information
    Establishment
    Transfer
    Termination

## Slide 9
High Level Protocol
This is the Structure of the messages, however let us make a few points about the High Level Protocol
  LIS2 – A is the Document that controls, (ASTM E 1394-91)
  Specifies message content for transferring information
  Structures the message exchange
  Provides common conventions for exchange of information
  Upload – data transmitted from an instrument to the LIS
  Download – data transmitted from LIS to instrument

## Slide 10
LIS Message
What are some of the things you will see in a typical message?
  <STX> = Start of Text Transmission character
  FN = Frame Number, (0 – 7)
  Text = The data content of the message
  <ETB> = End of Transmission Block, (Intermediate Frames)
  <ETX> = End of Text Transmission, (End Frames)
  Checksum = A 2 digit number, check for message integrity
  <CR> = Carriage Return, also a record delimiter, end of message

## Slide 11
LIS Message
What are some of the things you will see in a typical message?
  <LF> = Line Feed
  <ACK> = Acknowledge last frame received, ready for next frame
  <NAK> = No acknowledge of last frame, ready for resend
  <EOT> = End of Text, receiver ready but request sender to stop
  <ENQ> = Enquiry, ready to receive message
  <Ι> = Field Delimiter, separates adjacent fields
  <\> = Repeat Delimiter, separates equal parts of a field
  <^> = Caret, separates unequal parts of a field

## Slide 12
Message Structure
How is the typical message structured?
  (H) = Header Record, is sent once at the start of the message
  (P) = Patient Record, may be sent multiple times
  (O) = Order Record, must be preceded by the Patient Record
  (R) = Result Record, must be preceded by the Order Record
  (C) = Comment Record, additional text for explanation
  The LIS2-A document will set the structure of the message and the order that it must be sent and received
  It is up to the LIS Vendor to set up their program correctly to the parameters as set in the LIS2-A document
  Our instruments and software currently conform to the LIS2-A documents

## Slide 13
Download: LIS to BTA
H|\^& | | |LIS | | | | | BACT/ALERT | | P | 1 | 20021120112515<CR>
1 2 3 4 5 6 7 8 9 10 11 12 13 14
P | 1 | 245-13-3672 | | |SMITH^MARY | |19550401 | F | | | | |
2 3 4 5 6 7 8 9 10 11 12 13
0138^JONES | | | | | | | | | |20021101 | |ER^EMERGENCY R00M <CR>
14 15 16 17 18 19 20 21 22 23 24 25 26
The Black is the Message components
The ‘Ι’, or delimiter, is the space between frames
The Red is the Field Number of the message
( Normally will not be shown in the message, shown here to identify fields )
O | 1 | | 923240189 |^^^BC^BSA^SAQPTF7K^5\^^^BC^BSN^SNQRD8H8^5 |
1 2 3 4 5
This is the Universal
Test ID
This is the Carriage Return
Three Examples:

## Slide 14
Upload: BTA to LIS
H|\^& | | | BACT/ALERT^A.00 | | | | | | | P | 1 | 20021120112423 <CR>
1 2 3 4 5 6 7 8 9 10 11 12 13 14
P | 1 | 245-13-3672 <CR>
1 2 3
R|1 |^^^BC^BSA^!AQHTDR7 |–| | | | |F | | |20021113095221 |
2 3 4 5 6 7 8 9 10 11 12
20021118095221 |1B08 <CR>
13 14
H = Header; P = Patient Record; O = Order Record; R = Result Record; C = Comment
This is the Universal
Test ID
This is the Carriage Return
O| 1 | |923240190 | | | | | | | | | | | | | | | | | | | | | |F <CR>
1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 23 24 25 26
Sending Results from BTA to the LIS - Four Examples:

## Slide 15
Responsibility
Customer
  BacT/ALERT 3D User Configuration settings
  Cable to connect the LIS to BacT/ALERT 3D
  LIS Setup Parameters
    Baud Rate
    Parity
    Stop Bits
    Data Bits
Service Engineer
  Cable Information (Pin Layout)
  Setup User & LIS configuration settings on system
  Test the link

## Slide 16
RS-232C 9 Pin Cable
Pin #
Signal
1
Receive Line Signal Data
2
Receive Data
3
Transmit Data
4
Data Terminal Ready
5
Signal Ground
6
Data Set Ready
7
Request to Send
8
Clear to Send
9
Ring Indicator

## Slide 17
RS-232C 25 Pin Cable
Pin #
Signal
Pin #
Signal
1
Not Connected
14
Not Connected
2
Transmit Data
15
Not Connected
3
Receive Data
16
Not Connected
4
Request to Send
17
Not Connected
5
Clear to Send
18
Not Connected
6
Data Set Ready
19
Not Connected
7
Signal Ground
20
Data Terminal Ready
8
Receive Line Signal Data
21
Not Connected
9
Not Connected
22
Ring Indicator
10
Not Connected
23
Not Connected
11
Not Connected
24
Not Connected
12
Not Connected
25
Not Connected
13
Not Connected

## Slide 18
Record LIS Message
Diagnostic Password 12123434
Select 9.0 LIS on the left column
Select Message Log
Press the Right Arrow to display the message log for the LIS
Print or save this data to a disk
Send or Fax the information